[WS 1 in]

SPECTRUM CALCULATION ROUTINE

Version 2.0

[WS 4 in]

David E. Hess

Fluid Flow Group

Process Measurements Division

Chemical Science and Technology Laboratory

National Institute of Standards and Technology

Email: HESS@ENH.NIST.GOV

Phone: (301) 975-5937

[WS 0.5 in]

[April 1992]

New Features

*        Two additional window functions are available for tapering the data.

The first, referred to as the Minimum Sidelobe window, has a maximum

sidelobe level of -93 dB; the sidelobes then decay at a rate of 18

dB/octave.  The second window, the Maximum Decay window, has a maximum

sidelobe level of -61 dB, but these sidelobes decay at a rate of 42

dB/octave.

*        The structure of the data files to be processed has been modified to

allow inclusion of information about the voltage transformation factor

for each channel of data.  The file format has been changed to BINARY

in order to permit these data files to be written by programming

languages other than FORTRAN.

*        Memory for data arrays is now allocated dynamically at run-time, and

only the memory needed for the current file being processed is

allocated at any particular time.  Storage for these arrays can be

quite large and dynamic allocation of memory reduces the amount of

storage required within the SPECTRUM executable file.  This

modification has reduced the size of the executable file by 75%.

*        All internal temporary files created by the SPECTRUM routine are now

written in BINARY format.  This saves an additional small amount of

memory at run-time.

*        When an input data file consisted of only one channel of data to be

processed, the first version of SPECTRUM required that this data file

have an even number of records.  This restriction has been removed.

*        The manual included with the first version of SPECTRUM provides a

detailed account of the manner in which the amount of space on the

hard disk (or Ram disk) required by the routine for storage of

temporary files may be estimated.  Version 2 of SPECTRUM now

calculates at run-time the amount of temporary space required to

process the current file, and then checks the hard disk (or Ram disk)

to determine if that space is available.

*        The code has been restructured by removing portions of the code from

the Main module and placing these segments into functions or

subroutines.  The subsequent reduction in size of the Main module has

permitted additional optimization procedures contained within the

Microsoft Fortran Optimizing Compiler, V 5.1 to be employed.  Also,

the code for the Main module is now more readable.

*        All of the utility files included on the diskette have been modified

to use the new input data file format.  Each of these routines should

be used only with SPECTRUM V 2.0.

*        The SPECTRUM routine requires a computer equipped with an 80286 or

better microprocessor.  An 80287 or better numeric coprocessor is

recommended but is no longer required.

All information contained within this document supersedes the information

contained within the manual which refers to version 1.0 of the SPECTRUM

routine.  Any information found in the manual which is not countermanded by

statements made here is considered to still apply.

Diskette Contents

The following files are contained on the diskette.

SPECTRUM     FOR    66165  04-22-92  8:16a     Main program source code.

SPECTRUM     FI      1852  04-22-92  7:59a     Interface statements.

SPECTRUM     FN      2406  04-22-92  8:04a     Data declaration & common areas.

SPECTRUM     CFG     1197  04-23-92  3:50p     Optional setup file.

SPECTRUM     EXE   111592  04-23-92  3:03p     Works with or w/o a coprocessor.

SPEC387      EXE   103800  04-23-92  2:43p     Works only with a coprocessor.

COLOR        FOR     8314  04-22-92  8:05a     Source for noise generator.

COLOR        EXE    57452  04-23-92  3:07p     Executable for noise generator.

COLOR2       FOR     9073  04-23-92  3:13p     Source for 2-chan noise gen.

COLOR2       EXE    57872  04-23-92  3:13p     Executable for 2-chan noise gen.

SINE         FOR     4371  04-22-92  8:06a     Source for sinewave generator.

SINE         EXE    50490  04-23-92  3:09p     Executable for sinewave gen.

MAKDAT       FOR     8960  04-22-92  8:06a     Ascii to Binary converter

MAKDAT       EXE    51564  04-23-92  3:11p     source.

MAKCOEF      FOR     6296  04-22-92  8:07a     Asc to Bin converter executable.

MAKCOEF      EXE    44388  04-23-92  3:11p     Coefficient file creation

README       1ST     1594  04-28-92  2:54p     source.

Coeff. file creation executable.

Cover letter.

The README.1ST file is simply a reminder that this is the second version

of SPECTRUM and that an update to the manual exists and should be read prior to

the first use of the program.  All files with an .FOR, .FI or .FN extension are

Ascii files containing FORTRAN source code for the respective program.  These

files are not required when using SPECTRUM and need not be kept on the hard

disk.  If the SPECTRUM routine is recompiled, then three files are required

during compilation: SPECTRUM.FOR, SPECTRUM.FI and SPECTRUM.FN.  The SPECTRUM.FI

file contains interface statements for the functions and subroutines used in

SPECTRUM, and the SPECTRUM.FN file contains data declaration statements and

COMMON statements which permit data areas to be shared by both the Main module

and other program modules.  These two files are compiled along with the main

program file using the INCLUDE statement in appropriate locations in the main

program source code.  This program was compiled using the Microsoft FORTRAN

Optimizing Compiler, V 5.1.  The command used to compile the code and to

generate the SPECTRUM.EXE file, which may be used with or without a

coprocessor, is given by

FL /G2 /W2 /FPi SPECTRUM.FOR

In addition, an executable version of SPECTRUM has been provided which will

work only with a numeric coprocessor installed (80287, 80387 or 80486).  This

executable version is called SPEC387.EXE and consists of more streamlined code

which executes somewhat faster.  This version was compiled using the command

FL /G2 /W2 /FPi87 SPECTRUM.FOR

The use of the SPECTRUM.CFG file is optional and is described in the manual

which accompanied version 1 of SPECTRUM.  The COLOR, COLOR2 and SINE utility

routines generate data with known spectral densities; the data files produced

by these routines are compatible with SPECTRUM and may be used for testing the

operation of the program.  The use of these utility routines is reviewed in

detail in the version 1 manual.  Finally, the MAKDAT and MAKCOEF routines are

used to convert Ascii data files into the BINARY data files required by the

SPECTRUM routine.  These utilities are also explained in the version 1 manual.

Additional Windows

Two additional window functions for tapering the input data are available

in SPECTRUM V 2.0.  These windows have been added as possible alternatives to

the more commonly used windows made available in V 1.0.  The choice of a window

function depends upon such factors as the highest sidelobe level (dB), the rate

of decay of the sidelobes (dB/octave) and the halfwidth of the mainlobe (f T).

In general, as the width of the mainlobe is decreased, thereby increasing the

resolving power, the amplitude of the sidelobes is increased which intensifies

the problem of leakage.  In many cases it is useful to allow the width of the

mainlobe to increase somewhat in order to decrease the height of the sidelobes

and / or to reduce the rate of decrease in the sidelobe level.

With this end in mind, two windows were selected from Nuttall (1981) which

are highly optimized members of a family of windows discussed by Harris

(1978).  The first window, which will be denoted here as the Minimum Sidelobe

window, is the 4-term Blackman window with continuous first derivative.  This

window minimizes the maximum height of the sidelobes and also has a reasonably

good rate of decay of the sidelobes at the expense of a widened mainlobe.  The

second window will be referred to here as the Maximum Decay window; this is the

4-term Blackman window with continuous fifth derivative which maximizes the

rate of decay of the sidelobes and which also has a low maximum sidelobe

amplitude.  As expected, the mainlobe for this window is also wider than any of

the windows contained in the first version.

These windows are defined by

[EQN "center #^ w (t) #3#3 = #3#3 #( { #3#3 stack left [[ a sub 0 #3#3 - #3#3 a

where the coefficients for the Minimum Sidelobe window are given by

[EQN "center a sub 0 #3#3 = #3#3 0.355768 #4#4 a sub 1 #3#3 = #3#3 0.487396 #4#4

and those for the Maximum Decay window by

[EQN "center a sub 0 #3#3 = #3#3 10 over 32 #4#4 a sub 1 #3#3 = #3#3 15 over 32

These functions are plotted in figure 1 with the Hanning window shown for

comparison.  Notice that the Hanning window is a member of this family of

windows where the first two coefficients are 1/2 and the remaining coefficients

are zero.[PB]

The Fourier transforms for these windows are found to be

[EQN "center W #1 (f,T) #3#3 = #3#3 { #3#3 a sub 0 #3#3 #[ [sin #3 pi f T] over

where  [EQN "f sub 1 #3 = #3 1 #1 / #1 T"] .  The moduli of the Fourier

transforms for each of the above windows as well as for the Hanning window,

shown for comparison, are plotted in figures 2, 3 and 4.  In order to clearly

reveal the maximum sidelobe height and the rate of decay of these sidelobes,

the vertical axes of the plots are magnified by using a logarithmic scale; the

power response of the windows normalized by the peak response at

[EQN "f #3 = #3 0"]  and converted to decibels is defined by

[EQN "center #R[dB] #3#3 = #3#3 10 #3 #R[log] #3 [| W(f) #1 / #1 W(0)|] super 2"

The plots clearly show the optimal features for each window function.  A list

of the performance parameters for all of the windows contained within SPECTRUM

is included below for completeness.

Maximum        Maximum      Halfwidth of

Window         Sidelobe        Decay         Mainlobe

Type            (dB)       (dB/octave)       (f T)

Square           -13             6             1.0

Tukey           -13             18            1.1

Hanning          -31             18            2.0

Parzen           -27             12            2.0

Welch           -21             12            1.4

Min. Sidelobe       -93             18            4.0

Max. Decay         -61             42            4.0

Table 1  Window performance parameters

[PIC FIG1.CGM]

Figure 1  Additional window functions for tapering the data

[WS 0.5 in]

[PIC FIG2.CGM]

Figure 2  Power response for Hanning window[PB]

[PIC FIG3.CGM]

Figure 3  Power response for minimum sidelobe window

[WS 0.5 in]

[PIC FIG4.CGM]

Figure 4  Power response for maximum decay window

Input Data File Structure

The author has obtained a new A/D converter which has the capability of

altering the input voltage range to be digitized.  The selection of an input

voltage range may be made via software as opposed to changing jumpers on the

board itself.  After data has been sampled by the A/D converter, the author

has, in the past, written certain information to the first record of each data

file which describes the digitized information that follows.  With the added

capability offered by the new A/D converter, it was felt that the input voltage

range used by each channel in the input data file would be prudent pieces of

information to add to the file header.  The A/D converter has the ability to

convert up to eight differential inputs sequentially, and each channel may use

a different input voltage range.  Therefore, data files from this A/D converter

may contain up to eight channels of digitized information, and the second

record of the data file is now used to record the input voltage range used by

each channel.  These input data files are then used by a variety of routines

written by the author to perform various aspects of random data analysis.  The

SPECTRUM program is just one member of this family of routines.  Since the

SPECTRUM program can only process one or two channels of data, placing

information on the voltage range used by up to eight channels of data leads to

unnecessary information being included for six channels which do not exist.  As

silly as this may seem, it turns out to be useful for the author to adopt a

single universal input data file structure which can then be used by all of the

routines which examine this data file.  Therefore, this input data file

structure, while not optimal for the SPECTRUM program, is the most useful

universal setup for the author at this time.

All input data files must be sequential BINARY files.  These files contain

no internal formatting information, and thus, are the most useful type of files

to use when dealing with large quantities of data.  Sequential unformatted

FORTRAN data files were used as the file type for input data files in the first

version of SPECTRUM.  The drawback with unformatted FORTRAN data files is that

they may be written only by FORTRAN programs.  Using BINARY data files allows

the data to be written using other programming languages.  Furthermore,

unformatted FORTRAN data files do, in fact, use a small amount of internal

formatting information, whereas BINARY files do not.  Thus, BINARY data files

are slightly smaller than unformatted FORTRAN data files.  The first two

records of the data file contain descriptive information about the digitized

data to follow and form the file header.  Following the file header is the

actual digitized data.  The first record of the file header is changed only

slightly from version 1.0 of the SPECTRUM routine, and the text describing this

first record is taken from the manual issued with the first version.  Changes

are noted in boldface text as required.

The first record of the data file consists of a header containing four

quantities which characterize the data which follows. The first quantity is a

2-byte integer which gives the number of channels of data per record (1 or 2).

The next number is the record size in bytes and must be a 4-byte integer.  The

record size is twice the number of points per record if the data consists of

2-byte integer values, or the record size is four times the number of points

per record if the data consists of 4-byte real data.  The third quantity is a

2-byte integer which gives the number of records of data which follow, and the

last quantity is the time interval in microseconds between data points of the

same channel expressed as a 4-byte integer.  Recall that the largest positive

signed number that may be stored as a 2-byte integer is 32,767 and, for a

4-byte integer, the value is 2,147,483,647.

For example, suppose the data file consists of 100 records of one channel

of data saved as 2-byte integers and sampled at 10000 samples per second with

2048 data points per record.  The first record of the data file would contain

the four values:

[EQN "center 1 #4#4 4096 #4#4 100 #4#4 100 #4#4"]

Now suppose that two channels of data were sampled with a time interval of

0.016383 seconds between each successive data point resulting in a time

interval of 0.032766 seconds between each successive data point of the same

channel.  Furthermore, 40 records of data were sampled with 8192 data points

per record such that each channel of data consisted of 4096 data points.

Finally, the data were transformed into desired quantities and saved as 4-byte

floating-point numbers.  The first record for this data file should consist of

the following four numbers:

[EQN "center 2 #4#4 32768 #4#4 40 #4#4 32766"]

The second record of the data file consists of eight 2-byte integer values

which encode the input voltage range used by channels 1 to 8.  Acceptable

values are

[EQN "center matrix [[[# ] [# ] [# ] [# ] [# ] [# ] [#R[Value]] [# ] [#R[Range]]

The user should write a 2-byte integer value from 0 to 3 to indicate the

voltage range used by the first channel and the second channel.  Six additional

zero values must be entered for channels 3 through 8 which are not used by the

SPECTRUM routine.  For example, to indicate that the input voltage range for

the first channel was -5V to 5V and for the second channel was -1.25V to 1.25V,

the user would write the following to the second record of the data file.

[EQN "center 1 #3#3 3 #3#3 0 #3#3 0 #3#3 0 #3#3 0 #3#3 0 #3#3 0 #3#3"]

When the SPECTRUM program is run, the routine will ask if an alternate voltage

transformation factor should be used.  If the user responds with yes, then the

user must type in the voltage per digital level conversion factor to be used

for that channel, and this factor will be used in place of the factor stored in

record 2 of the data file.  After reading in the data file, the SPECTRUM

routine will convert the digitized data back into voltages using the conversion

factors shown above.  If none of the above voltage ranges are appropriate for a

particular data file, then the user should proceed as follows.  Store eight

zero values (2-byte integers) on the second record of the data file.  Then,

execute the SPECTRUM routine and enter the appropriate alternate conversion

factor when requested.

Following the first two records containing the file header is the actual

data.  When the data comes directly from an analog to digital converter

(12-bit, say), then the values range from -2048 to 2047 (or 0 to 4095) and are

appropriately stored as 2-byte integers.  Alternatively, if some preprocessing

is typically performed on the data, the data may be stored as 4-byte

floating-point numbers.  As a result of the Fast Fourier Transform (FFT)

algorithm used in this routine, the number of data points per channel per

record must be a power of 2.

In the first version of SPECTRUM, the program required an even number of

records of data in the input data file when the file consisted of one channel

of sampled data.  This is no longer required.  When the program submits the

data to the FFT module to be transformed, two records are submitted to the FFT

algorithm to be transformed simultaneously as a time-saving measure.  If an odd

number of records is detected in the input data file in the case of one channel

of data, then the program automatically appends an additional record of data

containing a copy of the data values of the previous record.  After

transformation, the results of the added record are discarded so that the

output is not corrupted by the internal addition of a record of data.  When two

channels of data are sampled, an even or odd number of records is not a factor

because one record from each channel is submitted to the FFT module for

simultaneous transformation.

One or two separate time series (channels) may be stored per record.  The

method of storing the data values in each record follows the standard procedure

used by analog to digital converters:

[EQN "center #R[one] #4 #R[channel] #^ #R[:] #4#4 x sub 0 #3 x sub 1 #3 x sub 2

where the notation is as defined above.  This convention must be followed even

if the data are processed and then stored as floating-point numbers.  When two

channels of data are sampled and then stored to a data file, the values for

each channel of data should be of the same data type: either 2-byte integers or

4-byte floating-point numbers.  If the user is unfamiliar with the task of

creating BINARY data files, then the user may instead create these files in an

ASCII file format.  A utility MAKDAT which is included on the diskette can then

be used to write the file header and transform the ASCII data files into the

BINARY data files required by the spectrum routine.  See the section on utility

files in the manual included with version 1.0.

This section gives a few lines of code which are representative of the

manner in which these data files are written and read.  When writing, the file

is first opened, the file header is written, the data is written and then the

file is closed.

INTEGER*2 ICHANS, NUMREC

INTEGER*4 IRSIZE, IDELTMS, N, I

INTEGER*2 NDATA (NMAX), GAIN (0 : 7)

OPEN (2, FILE = 'A001.DAT', STATUS = 'UNKNOWN',

ACCESS = 'SEQUENTIAL', FORM = 'BINARY')

WRITE (2) ICHANS, IRSIZE, NUMREC, IDELTMS

WRITE (2) (GAIN (I), I = 0, 7)

DO J = 1, NUMREC

... put data record into NDATA array ...

WRITE (2) (NDATA (I), I = 1, N)

... get next record of data ...

ENDDO

CLOSE (2, STATUS = 'KEEP')

The same procedure is followed when the data file is opened and read by

the spectrum routine.  The file is opened, the file header is read, the data is

read and then the file is closed.

INTEGER*2 ICHANS, NUMREC

INTEGER*4 IRSIZE, IDELTMS, N, I

INTEGER*2 NDATA (NMAX), GAIN (0 : 7)

OPEN (2, FILE = 'A001.DAT', STATUS = 'OLD',

ACCESS = 'SEQUENTIAL', FORM = 'BINARY')

READ (2) ICHANS, IRSIZE, NUMREC, IDELTMS

READ (2) (GAIN (I), I = 0, 7)

N = IRSIZE/2

DELT = FLOAT (IDELTMS) / 1000000.0

DO J = 1, NUMREC

READ (2) (NDATA (I), I = 1, N)

... process this record of data ...

ENDDO

CLOSE (2, STATUS = 'KEEP')

MAKCOEF Utility Routine

The use of this utility routine for the generation of coefficient data

files for use with SPECTRUM is essentially the same.  The output of the program

is now a BINARY coefficient data file vis a vis the unformatted FORTRAN

coefficient data file generated in the first version.  The SPECTRUM routine now

allocates memory dynamically for all internal arrays at run-time.  This

improvement in the main program has unfortunately introduced an additional

restriction on the generation of coefficient data files.  The first set of

coefficients contained in the coefficient data file must correspond to the

first file to be processed by SPECTRUM, and the last set of coefficients

contained in the coefficient data file must correspond to the last file to be

processed by SPECTRUM.  The total number of sets of coefficients contained in

the coefficient data file should be given by

[EQN "center #R[numsets] #3#3 = #3#3 #R[last] #3#3 - #3#3 #R[first] #3#3 + #3#3

The total number of files processed by SPECTRUM need not match the total number

of sets of coefficients contained in the coefficient data file as long as the

first and last sets of coefficients correspond to the first and last files

processed, and the remainder of the files to be processed correspond to sets of

coefficients that are contained between the first and last sets in the

coefficient data file.  This restriction will be clarified below using the

example given on pp 70-71 from the manual included with the first version.

Suppose that it was desired to process the data files D061.DAT thru

D086.DAT in a batch by the spectrum routine.  Each of these data files

consisted of two channels of 2-byte integers taken from an experiment in fluid

mechanics, say.  Furthermore, suppose that the quantities sampled were two

components of a fluid velocity, which were transduced into voltages by some

means and then sampled by an A/D converter.  The user knows the calibration

data for the transducer, and the required transformation from voltages back

into velocities can be accomplished for each velocity component by using a

polynomial of up to fourth degree in the form:

[EQN "center u sub n #3#3 = #3#3 b sub 0 #3 + #3 b sub 1 v sub n #3 + #3 b sub 2

where  [EQN "v sub n"]  is the nth data point expressed as a voltage,

[EQN "b sub 0 #1 , #3 b sub 1 #1 , ... , b sub 4"]  are the coefficients of the

velocity transformation and  [EQN "u sub n"]  is the nth data point which has

been converted to a velocity.  Now, suppose further that the data in each of

the files were taken over a period of time in which the calibration

coefficients changed; thus, a different set of coefficients are required for

each of the 26 files to be processed.  Following the instructions in the

section on coefficient data files, this scenario would require the creation of

two BINARY coefficient data files, DCON.DAT and DCON2.DAT, each containing the

26 sets of five coefficients  [EQN "b sub 0 #1 , #3 b sub 1 #1 , ... , b sub 4"]

for the velocity transformation for channel 1 and channel 2, respectively.

The user should create each ASCII data file with a four character name

followed by the extension [.ASC].  The names should conform to the various

naming conventions specified earlier.  The ASCII file must contain numbers

only, although blank lines may be inserted into the file if desired.  The

numbers will be transferred to 4-byte floating point quantities in the

unformatted data file thereby limiting the coefficients to single precision (6

or 7 decimal places).  The coefficients may be stored one per line or many per

line in the ASCII file.  The first number in the file will be  [EQN "b sub 0"]

for channel 1 of data file D061.DAT; the fifth number in the file should be

[EQN "b sub 4"]  for channel 1 of data file D061.DAT; the sixth number in the

file will be  [EQN "b sub 0"]  for channel 1 of data file D062.DAT, and so on.

The 130th and last number in the file should be  [EQN "b sub 4"]  for channel 1

of data file D086.DAT.  The user should then create a second ASCII file in a

similar manner specifying all of the coefficients for channel 2 of the 26 files

to be processed.  The utility MAKCOEF should then be executed twice, first

transferring the contents of the first ASCII file (for channel 1) to DCON.DAT

and then for the second ASCII file (for channel 2) to DCON2.DAT.  A session

using MAKCOEF would proceed as follows.

makcoef

Enter first letter of data file to

which these coefficients will be associated : d

Are these coefficients for channel (1 or 2) : 1

Enter # of sets of coefficients : 26

Enter starting file number : 61

Enter ASCII input file name (4 chars) : chn1

C O E F F I C I E N T   F I L E   C R E A T I O N   U T I L I T Y

Creating DCON.DAT now.

# sets of coeffs = 26

# coeffs in each set = 5

# of starting file = 61

Program terminated successfully.

The MAKCOEF routine will read the ASCII file CHN1.ASC and will write the

following file header to the BINARY file DCON.DAT:

[EQN "center 26 #4#4 61"]

The coefficients for channel 1 are then transferred to the BINARY file.  Note

that 5 coefficients per set are required; therefore, this number is not written

to the file header.  If, at a later time, the user wishes to reprocess the

files D074.DAT through D080.DAT (perhaps specifying different options in the

spectrum routine), it is necessary to re-create the two coefficient data files;

the new data files will contain 7 sets of coefficients beginning with those

corresponding to D074.DAT and ending with those corresponding to D080.DAT.  If

the user wishes to process the data files in an arbitrarily numbered order (not

consecutively), the coefficient data files may need to be changed.  For

example, if the files D064.DAT, D069.DAT and D086.DAT were to be processed by

SPECTRUM, this would necessitate a change in the coefficient data files

requiring 23 sets of coefficients in each file: the first set corresponding to

D064.DAT and the last set to D086.DAT.  On the other hand, the files D061.DAT,

D064.DAT, D069.DAT and D086.DAT could be processed by SPECTRUM without changing

the coefficient data files since the first and last files correspond to the

first and last sets of coefficients.  The converse is also true, however.  If

the user desires at the outset to process only the four arbitrarily ordered

files in the latter case above, two coefficient data files must be created

which contain not 4, but 26 sets of 5 coefficients for the data files D061.DAT

through D086.DAT.  The spectrum routine requires that sets of coefficients in

the coefficient data file be listed for consecutively numbered files even if

not all of the sets of coefficients are used.  With this in mind, it is perhaps

best to store all of the coefficients for an entire set of input data files in

a master ASCII file; then members of the set can be processed by the spectrum

routine either consecutively or in an arbitrarily numbered order at any

subsequent time and the BINARY coefficient data files can be easily created by

editing the master ASCII file and using MAKCOEF.  Of course, the user could

transform the voltages back into dimensional quantities using his/her own code

and store the resulting floating-point numbers into ASCII data files (then use

MAKDAT) or directly into BINARY data files to be processed by SPECTRUM; then,

all of this mess concerning coefficient data files could be avoided altogether.

*References

Harris, F.J. 1978 On the use of windows for harmonic analysis with the discrete

Fourier transform, Proceedings of the IEEE, 66, 51-83.

Nuttall, A.H. 1981 Some windows with very good sidelobe behavior, IEEE

Transactions on Acoustics, Speech and Signal Processing, ASSP-29, 84-91.